## Slide 2
EBSL1-ID18
Ray tracing and wave optics study of WBDM
Juan Reyes Herrera
Manuel Sanchez del Rio
AAPU/MEG
11/08/2021 (version 1)
18/11/2021 (version 2)

## Slide 3: OASYS models using SHADOW (ray tracing) and WOFRY (wave optics)
Configurations: 7 keV, WBDM Si @ 3 mrad, coherence slit: 40 x 206.8 mm, f1 (h: 41.69 m, v: 49.73 m), f2 (h: 26.15 m, v:39.72 m)
Slope errors were considered for both WB mirrors in both SHADOW and WOFRY
SHADOW used hybrid mode for the slit only (to get reasonable vertical sizes))
Shadow (ray tracing)
Wofry (wave optics)
https://raw.githubusercontent.com/srio/paper-transfocators-resources/main/workspaces/Hybrid-Shadow/EBS_ID18_shadow_WBDM.ows
https://raw.githubusercontent.com/srio/paper-transfocators-resources/main/workspaces/EBS_ID18_wofry_WBDM.ows

## Slide 4: Profile errors of WBDM
length [m]
Original central profile*
*From file provided by Raymond Barrett (ring256_TypbeB_F127001_frontside_ontissue_meas2__avg_2D.txt)
profile [m]
Best fitting curve
Detrending and Gaussian smoothing

## Slide 5: Profile errors of WBDM
After applying the surface modification to the 2D surface:
By simply scaling the profiles, simulation were performed for 4 conditions:
Ideal surfaces
mod. surface: 165 nrad
mod. Surface to get about ½ slope error*
mod. Surface to get about 2x slope error
* Because non linearity in Gaussian filtering, this is obtained multiplying height by 0.35

## Slide 6: Beam profile – Coherence Slit (36 m) SHADOW
~80 nrad
ideal
~165 nrad
~330 nrad
z [mm]
z [mm]
z [mm]
z [mm]

## Slide 7: Beam profile – Coherence Slit (36 m) WOFRY
~80 nrad*
ideal
~165 nrad*
~330 nrad*
z [mm]
z [mm]
z [mm]
z [mm]
*The horizontal intensity has an abrupt cut of the tails: this is an artifact due to the propagation of the high modes, and should be improved

## Slide 8: Beam profile – Transfocator 1 entrance (65 m) SHADOW
~80 nrad
ideal
~165 nrad
~330 nrad
z [mm]
z [mm]
z [mm]
z [mm]

## Slide 9: Beam profile – Transfocator 1 entrance (65 m) WOFRY
~80 nrad
ideal
~165 nrad
~330 nrad
z [mm]
z [mm]
z [mm]
z [mm]

## Slide 10: Beam profile – Sample position (200 m) SHADOW
~80 nrad
ideal
~165 nrad
~330 nrad
z [mm]
z [mm]
z [mm]
z [mm]

## Slide 11: Beam profile – Sample position (200 m) WOFRY
~80 nrad
ideal
~165 nrad
~330 nrad
z [mm]
z [mm]
z [mm]
z [mm]